Laser Report

AFRL 711HPW/RHDO 2.5.3.64

LHAZ Plugin 5.2.3.2

LTMC Version 3.2.2.7 / Adapter 3.1.0.19

Tuesday, November 24, 2015

Laser Name:				Spectral Physics 162A07

Laser Parameters:

Wavelength:			457.9 nm

Output Mode:			ContinuousWave

Average Power:		9 mW

Beam Profile:			Circular

Beam Distribution:		Gaussian

Beam Divergence:		0.5 mrad

Beam Waist Diameter:		0.5 cm

Beam Waist Range:		0

MPE Computations:

Exposure Duration:		0.25 s

Exposure Range:		10 cm

MPE (Eye):			2.598e-003 W/cm²

Limiting Aperture (Eye):	0.7 cm

Class 1 AEL (Eye):		1.000e-003 W

Limiting Aperture (Skin):	0.35 cm

MPE (Skin):			2.000e-001 W/cm²

Lower exposure MPE values are required for visible wavelength lasers when the eye is immobilized or has a large pupil such as health care with ophthalmic instruments or in research situations.  See ANSI Z136.1, Section 8.3.

Classification:				Class 3B

Description:

11265 or 12159

Hazard Distances and OD Requirements:

Ocular	(10 cm, Unaided Viewing, Existing OD = 0)

Exposure Duration:		0.25 s

NOHD:				39 m

At Viewing Distance:		10 cm

Maximum OD:			1.0

At Range OD:			0.9

Skin (10 cm, Existing OD = 0)

Exposure Duration:		600 s

NSHD:				0 m

At Exposure Distance:		10 cm

Maximum OD:			0.0

At Range OD:			0.0

Diffuse Reflection Hazard Analysis:

Laser to Target Range:	3 m

Target Reflectance:		10.00 %

Viewing Angle:		0 deg

Ocular	Hazards

Exposure Duration:		600 s

NHZ:				0.0 m

At Viewing Distance:		3 m

OD Required:			0.0

Skin

Exposure Duration:		600 s

NHZ (Skin):			0.0 m

At Exposure Distance:		3 m

OD Required:			0.0

Viewing Conditions:

Atm. Attenuation Coeff:	2.2961e-06 cm-1 (1/cm)

Aided Viewing Used:		False

Optics Transmittance:		N/A

Optics Objective Diam:	N/A

Optics Exit Diam:		N/A